## CT_protocol_data
scan_protocol_no | scan_protocol_name | scan_series | Speed | Type | Rows | HiRes | Shuttle | Tilt | SFOV | kV | SmartmA | NoiseIndex | MinmA | MaxmA | Message | Timer | Light | CTDI | CTDI NV | DLP | DLP NV | mA | pitch
1.5 | QQ Brain-Head Routine TFI-H | Series 1 | 0.8 | Helical | 64 | false | false | S0.0 | Adult Head | 100 | true | 4 | 150 | 515 | 102 | 0 | 0 | 55.19 | 1000 | 1880.49 | 1000
1.3 | Dental non-Helical |  | 1 | Axial | SmartCollimation | true | false | S0.0 | Adult Head | 140 | true | 18 | 80 | 150 | 101 | 0 | 0 | 29.72 | 1000 | 475.46 | 1000
1.4 | QQ Brain non-Helical Routine |  | 1 | Axial | 256 | false | false | S0.0 | Adult Head | 100 | true | 4 | 200 | 720 | 101 | 0 | 0 | 66.38 | 1000 | 1062.1 | 1000
1.11 | Head CTA 1Volume SmartPrep Routine | Series 2 Group 1 Scan Settings | 1 | Axial | SmartCollimation | false | false | S0.0 | Adult Head | 100 | true | 3.2 | 200 | 720 | 101 | 0 | 0 | 55.26 | 1000 | 884.2 | 1000
1.11 | Head CTA 1Volume SmartPrep Routine | Series 3 Group 1 Scan Settings | 1 | Axial | SmartCollimation | true | false | S0.0 | Adult Head | 100 | true | 3.2 | 200 | 570 | 101 | 0 | 0 | 52.5 | 1000 | 839.99 | 1000
1.17 | Head CTA+CTV 1Volume SmartPrep | Series 1 Group 1 Recon 1 Recon Settings | 1 | Axial | SmartCollimation | false | false | S0.0 | Adult Head | 100 | true | 3.2 | 200 | 720 | 101 | 0 | 0 | 55.26 | 1000 | 884.2 | 1000
1.17 | Head CTA+CTV 1Volume SmartPrep | Series 2 Group 1 Scan Settings | 1 | Axial | SmartCollimation | false | false | S0.0 | Adult Head | 100 | true | 3.2 | 200 | 720 | 101 | 0 | 0 | 55.26 | 1000 | 884.2 | 1000
1.17 | Head CTA+CTV 1Volume SmartPrep | Series 3 Group 1 Scan Settings | 1 | Axial | SmartCollimation | true | false | S0.0 | Adult Head | 100 | true | 3.2 | 200 | 570 | 101 | 0 | 0 | 52.5 | 1000 | 839.99 | 1000
1.17 | Head CTA+CTV 1Volume SmartPrep | Series 3 Group 2 Scan Settings | 1 | Axial | SmartCollimation | true | false | S0.0 | Adult Head | 100 | true | 3.2 | 200 | 570 | 101 | 0 | 0 | 52.5 | 1000 | 839.99 | 1000
1.2 | Brain Helical Routine | Series 2 Group 1 Scan Settings | 0.8 | Helical | 64 | false | false | S0.0 | Adult Head | 100 | true | 3.2 | 150 | 515 | 101 | 0 | 0 | 55.14 | 1000 | 1216.97 | 1000
1.2 | Brain Helical Routine | Series 3 Group 1 Scan Settings | 0.8 | Helical | 64 | false | false | S0.0 | Adult Head | 100 | true | 3.2 | 150 | 515 | 101 | 0 | 0 | 55.14 | 1000 | 1327.24 | 1000
1.18 | Head CTA+CTV Helical SmartPrep | Series 2 Group 1 Scan Settings | 0.8 | Helical | 64 | false | false | S0.0 | Adult Head | 100 | true | 3.2 | 200 | 515 | 101 | 0 | 0 | 49.01 | 1000 | 885.7 | 1000
1.18 | Head CTA+CTV Helical SmartPrep | Series 3 Group 1 Scan Settings | 0.6 | Helical | 64 | true | false | S0.0 | Adult Head | 100 | true | 3.2 | 200 | 570 | 101 | 0 | 0 | 48.82 | 1000 | 881.4 | 1000
1.18 | Head CTA+CTV Helical SmartPrep | Series 3 Group 2 Scan Settings | 0.6 | Helical | 64 | true | false | S0.0 | Adult Head | 100 | true | 3.2 | 200 | 570 | 101 | 0 | 0 | 48.82 | 1000 | 881.4 | 1000
1.14 | Head-Neck CTA Helical SmartPrep Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Adult Head | 100 | false | 11.6 | 200 | 700 | 101 | 0 | 0 | 6.63 | 1000 | 218.13 | 1000
1.14 | Head-Neck CTA Helical SmartPrep Routine | Series 3 Group 1 Scan Settings | 0.6 | Helical | 64 | true | false | S0.0 | Adult Head | 100 | false | 8 | 200 | 570 | 101 | 0 | 0 | 34.33 | 1000 | 1129.16 | 1000
1.13 | Head Perfusion Routine | Series 2 Group 1 Scan Settings | 0.35 | Axial | SmartCollimation | false | false | S0.0 | Adult Head | 80 | 285 |  |  |  | 101 | 0 | 0 | 202.74 | 1000 | 3243.76 | 1000
1.19 | QQ (Plain) | Series 2 Group 1 Scan Settings | 1 | Axial | SmartCollimation | false | false | S0.0 | Adult Head | 100 | true | 3.2 | 200 | 720 | 101 | 0 | 0 | 55.26 | 1000 | 884.2 | 1000
1.19 | QQ (Plain) | Series 3 Group 1 Scan Settings | 0.35 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 200 | 700 | 102 | 0 | 2 | 3.88 | 1000 | 291.47 | 1000
1.2 | QQ (Plain + CE 2Phase) | Series 2 Group 1 Scan Settings | 1 | Axial | SmartCollimation | false | false | S0.0 | Adult Head | 100 | true | 3.2 | 200 | 720 | 101 | 0 | 0 | 55.26 | 1000 | 884.2 | 1000
1.2 | QQ (Plain + CE 2Phase) | Series 3 Group 1 Scan Settings | 0.35 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 200 | 700 | 102 | 0 | 2 | 3.88 | 1000 | 291.47 | 1000
1.2 | QQ (Plain + CE 2Phase) | Series 4 Group 1 Scan Settings | 0.35 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 200 | 700 | 102 | 0 | 2 | 3.88 | 1000 | 291.47 | 1000
1.21 | AI | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Adult Head | 100 | true | 2 | 350 | 700 | 101 | 0 | 0 | 76.58 | 1000 | 1383.91 | 1000
1.21 | AI | Series 3 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Medium | 100 | true | 6 | 300 | 720 | 102 | 0 | 2 | 14.03 | 1000 | 969.02 | 1000
1.22 | GSI - Head Helical | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Adult Head | 120 |  |  |  |  | 101 | 0 | 0 | 84.62 | 1000 | 1866.67 | 1000 | 395 | 33
1.22 | GSI - Head Helical | Series 3 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Adult Head | 120 |  |  |  |  | 101 | 0 | 0 | 84.62 | 1000 | 2035.9 | 1000 | 395 | 33
1.15 | GSIX - Head CTA Helical SmartPrep Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Adult Head | 120 |  |  |  |  | 101 | 0 | 0 | 35.65 | 1000 | 643.79 | 1000 | 280 | 33
1.15 | GSIX - Head CTA Helical SmartPrep Routine | Series 3 Group 1 Scan Settings | 0.5 | Helical | 64 | false | false | S0.0 | Adult Head | 120 |  |  |  |  | 101 | 0 | 0 | 30.08 | 1000 | 543.26 | 1000 | 280 | 33
1.23 | GSI - Neck CTA SmartPrep | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 101 | 0 | 0 | 8.8 | 1000 | 317.43 | 1000 | 185 | 127
1.23 | GSI - Neck CTA SmartPrep | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 101 | 0 | 0 | 8.8 | 1000 | 317.43 | 1000 | 185 | 127
1.1 | Brain 100mm 2Volume | Series 2 Group 1 Scan Settings | 1 | Axial | 160 | false | false | S0.0 | Adult Head | 100 | true | 3.2 | 200 | 720 | 101 | 0 | 0 | 76.05 | 1000 | 1368.94 | 1000
1.1 | Brain 100mm 2Volume | Series 3 Group 1 Scan Settings | 1 | Axial | 160 | false | false | S0.0 | Adult Head | 100 | true | 3.2 | 200 | 720 | 101 | 0 | 0 | 63.38 | 1000 | 1140.79 | 1000
1.16 | GSIX - Head-Neck CTA Helical SmartPrep Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Adult Head | 120 |  |  |  |  | 101 | 0 | 0 | 35.65 | 1000 | 643.79 | 1000 | 280 | 33
1.16 | GSIX - Head-Neck CTA Helical SmartPrep Routine | Series 3 Group 1 Scan Settings | 0.5 | Helical | 64 | false | false | S0.0 | Adult Head | 120 |  |  |  |  | 101 | 0 | 0 | 15.76 | 1000 | 518.42 | 1000 | 280 | 63
1.6 | Brain non-Helical Routine | Series 2 Group 1 Scan Settings | 1 | Axial | 256 | false | false | S0.0 | Adult Head | 100 | true | 4 | 200 | 720 | 101 | 0 | 0 | 66.38 | 1000 | 1062.1 | 1000
1.6 | Brain non-Helical Routine | Series 3 Group 1 Scan Settings | 1 | Axial | SmartCollimation | false | false | S0.0 | Adult Head | 100 | true | 4 | 200 | 720 | 101 | 0 | 0 | 55.32 | 1000 | 885.08 | 1000
1.9 | Dental Helical 140kV Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | true | false | S0.0 | Adult Head | 140 | true | 17 | 80 | 130 | 101 | 0 | 2 | 33.38 | 1000 | 602.74 | 1000
1.7 | Brain Head Routine TFI-H | Series 2 Group 1 Scan Settings | 0.8 | Helical | 64 | false | false | S0.0 | Adult Head | 100 | true | 4 | 150 | 515 | 102 | 0 | 0 | 55.19 | 1000 | 2432.42 | 1000
1.8 | GSIX - Head Helical | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Adult Head | 120 |  |  |  |  | 101 | 0 | 0 | 84.62 | 1000 | 1866.67 | 1000 | 395 | 33
1.8 | GSIX - Head Helical | Series 3 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Adult Head | 120 |  |  |  |  | 101 | 0 | 0 | 84.62 | 1000 | 2035.9 | 1000 | 395 | 33
1.1 | GSIX Dental Helical | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 140 |  |  |  |  | 101 | 0 | 2 | 8.8 | 1000 | 176.65 | 1000 | 185 | 127
1.12 | Head CTA+CTV 1Volume SmartPrep Routine | Series 2 Group 1 Scan Settings | 1 | Axial | SmartCollimation | false | false | S0.0 | Adult Head | 100 | true | 3.2 | 200 | 720 | 101 | 0 | 0 | 55.32 | 1000 | 885.08 | 1000
1.12 | Head CTA+CTV 1Volume SmartPrep Routine | Series 3 Group 1 Scan Settings | 1 | Axial | SmartCollimation | true | false | S0.0 | Adult Head | 100 | true | 3.2 | 200 | 570 | 101 | 0 | 0 | 52.55 | 1000 | 840.83 | 1000
1.25 | RT Head | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 120 | false | 6 | 150 | 515 | 102 | 0 | 0 | 39.48 | 1000 | 1739.87 | 1000
2.1 | Sinus&Face Axial | Series 2 Group 1 Scan Settings | 1 | Axial | SmartCollimation | false | false | S0.0 | Adult Head | 100 | true | 6.1 | 100 | 330 | 101 | 0 | 0 | 30.39 | 1000 | 486.31 | 1000
2.1 | Sinus&Face Axial | Series 3 Group 1 Scan Settings | 1 | Axial | SmartCollimation | false | false | S0.0 | Adult Head | 100 | true | 6.1 | 100 | 330 | 101 | 0 | 0 | 9.21 | 1000 | 147.37 | 1000
2.3 | Inner Ear non-Helical Routine | Series 2 Group 1 Scan Settings | 1 | Axial | SmartCollimation | true | false | S0.0 | Adult Head | 100 | true | 6.1 | 100 | 330 | 101 | 0 | 0 | 31.37 | 1000 | 313.72 | 1000
2.2 | QQ Orbit Sinus Helical Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Adult Head | 100 | true | 8 | 100 | 330 | 101 | 0 | 0 | 9.2 | 1000 | 166.24 | 1000
2.2 | QQ Orbit Sinus Helical Routine | Series 3 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Adult Head | 100 | true | 8 | 100 | 330 | 101 | 0 | 0 | 9.2 | 1000 | 166.24 | 1000
2.4 | Navigation Orbit Sinus Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Adult Head | 100 | true | 6.1 | 100 | 330 | 101 | 0 | 0 | 9.19 | 1000 | 166.07 | 1000
3.1 | Head-Neck | Series 2 Group 1 Scan Settings | 0.5 | Helical | 64 | false | false | S0.0 | Adult Head | 100 | true | 3.2 | 150 | 580 | 101 | 0 | 0 | 55.5 | 1000 | 1224.91 | 1000
3.1 | Head-Neck | Series 2 Group 2 Scan Settings | 0.5 | Helical | 64 | false | false | S0.0 | Adult Head | 100 | true | 5 | 150 | 700 | 101 | 0 | 0 | 66.6 | 1000 | 1569.79 | 1000
3.1 | Head-Neck | Series 3 Group 1 Scan Settings | 0.5 | Helical | 64 | false | false | S0.0 | Adult Head | 100 | true | 3.2 | 150 | 580 | 101 | 0 | 0 | 55.5 | 1000 | 1224.91 | 1000
3.1 | Head-Neck | Series 3 Group 2 Scan Settings | 0.5 | Helical | 64 | false | false | S0.0 | Adult Head | 100 | true | 5 | 150 | 700 | 101 | 0 | 0 | 66.6 | 1000 | 1569.79 | 1000
3.3 | Neck Helical Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Small Body | 100 | true | 5 | 150 | 700 | 102 | 0 | 2 | 17.19 | 1000 | 566.03 | 1000
3.3 | Neck Helical Routine | Series 3 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Small Body | 100 | true | 5 | 150 | 700 | 102 | 0 | 2 | 17.19 | 1000 | 737.95 | 1000
3.4 | Neck - Chest Helical Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 250 | 700 | 102 | 0 | 2 | 14.32 | 1000 | 803.77 | 1000
3.4 | Neck - Chest Helical Routine | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 250 | 700 | 102 | 0 | 2 | 14.32 | 1000 | 803.77 | 1000
3.5 | Neck - Pelvis Helical Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 250 | 700 | 102 | 0 | 2 | 14.32 | 1000 | 1161.84 | 1000
3.5 | Neck - Pelvis Helical Routine | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 250 | 700 | 102 | 0 | 2 | 14.32 | 1000 | 1161.84 | 1000
3.2 | C-Spine non Helical Routine | Series 2 Group 1 Scan Settings | 1 | Axial | SmartCollimation | false | false | S0.0 | Small Body | 100 | true | 5 | 150 | 700 | 101 | 0 | 2 | 9.36 | 1000 | 262.05 | 1000
3.8 | GSI - Neck | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.48 | 1000 | 279.1 | 1000 | 215 | 63
3.8 | GSI - Neck | Series 3 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.48 | 1000 | 363.92 | 1000 | 215 | 63
3.9 | GSI - C-Spine | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 101 | 0 | 0 | 8.48 | 1000 | 245.17 | 1000 | 215 | 63
3.1 | QQ C-Spine Helical MAR Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Small Body | 100 | true | 5 | 150 | 700 | 102 | 0 | 2 | 8.66 | 1000 | 260.45 | 1000
3.12 | Neck 4D | Series 2 Group 1 Scan Settings | 0.35 | Cine | 160 | false | false | S0.0 | Adult Head | 80 | true | 11.6 | 100 | 600 | 102 | 0 | 2 | 54.11 | 1000 | 541.12 | 1000
3.7 | GSIX C-Spine Routine | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 100 |  |  |  |  | 102 | 0 | 2 | 14.49 | 1000 | 476.64 | 1000 | 270 | 63
3.6 | Neck CTA Helical SmartPrep Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Adult Head | 100 | true | 11.6 | 200 | 700 | 102 | 0 | 1 | 6.62 | 1000 | 217.91 | 1000
3.6 | Neck CTA Helical SmartPrep Routine | Series 3 Group 1 Scan Settings | 0.6 | Helical | 64 | true | false | S0.0 | Adult Head | 100 | true | 11.6 | 200 | 570 | 102 | 0 | 1 | 34.3 | 1000 | 1128.03 | 1000
3.6 | Neck CTA Helical SmartPrep Routine | Series 3 Group 2 Scan Settings | 0.6 | Helical | 64 | true | false | S0.0 | Adult Head | 100 | true | 11.6 | 200 | 570 | 102 | 0 | 1 | 34.3 | 1000 | 1128.03 | 1000
3.11 | GSI - Neck - Copy | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 9.07 | 1000 | 298.49 | 1000 | 185 | 63
3.11 | GSI - Neck - Copy | Series 3 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 9.07 | 1000 | 389.2 | 1000 | 215 | 63
3.14 | GSIX Special | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Small Body | 100 | true | 5 | 150 | 700 | 102 | 0 | 2 | 17.19 | 1000 | 256.57 | 1000
3.14 | GSIX Special | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 704.58 | 1000 | 185 | 127
3.14 | GSIX Special | Series 4 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Small Body | 100 | true | 5 | 150 | 700 | 102 | 0 | 2 | 17.19 | 1000 | 256.57 | 1000
3.15 | GSIX neck-Pelvi CE Only | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 757.37 | 1000 | 185 | 127
3.13 | GSIX neck-Pelvi P+CE | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.1 | 1000 | 1781.8 | 1000
3.13 | GSIX neck-Pelvi P+CE | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 493.41 | 1000 | 185 | 127
4.12 | Hand Tendon 140keV nonHelical | Series 2 Group 1 Scan Settings | 1 | Axial | SmartCollimation | true | false | S0.0 | Small Body | 140 | true | 7.8 | 200 | 405 | 102 | 0 | 0 | 22.43 | 1000 | 448.67 | 1000
4.1 | Shoulder Axial | Series 2 Group 1 Scan Settings | 1 | Axial | SmartCollimation | true | false | S0.0 | Medium | 100 | true | 7.8 | 150 | 570 | 101 | 0 | 0 | 12.15 | 1000 | 291.65 | 1000
4.4 | QQ Arm on Abd Helical MAR | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Medium | 100 | true | 7.8 | 200 | 700 | 102 | 0 | 2 | 10.25 | 1000 | 267.29 | 1000
4.3 | Arm on Abd Axial | Series 2 Group 1 Scan Settings | 1 | Axial | SmartCollimation | false | false | S0.0 | Medium | 120 | true | 7.8 | 200 | 700 | 102 | 0 | 2 | 19.04 | 1000 | 456.86 | 1000
4.6 | QQ Arm over Head Helical MAR | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | true | false | S0.0 | Small Body | 100 | true | 7.8 | 100 | 450 | 102 | 0 | 0 | 8.65 | 1000 | 225.44 | 1000
4.5 | Arm over Head Axial | Series 2 Group 1 Scan Settings | 1 | Axial | SmartCollimation | true | false | S0.0 | Small Body | 100 | true | 7.8 | 100 | 450 | 102 | 0 | 0 | 9.79 | 1000 | 215.47 | 1000
4.8 | GSI - Shoulder | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 120 |  |  |  |  | 101 | 0 | 0 | 13.5 | 1000 | 336.21 | 1000 | 380 | 63
4.9 | GSI - Arm on Abd | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 120 |  |  |  |  | 102 | 0 | 2 | 13.5 | 1000 | 336.21 | 1000 | 380 | 63
4.1 | GSI - Arm over Head | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 120 |  |  |  |  | 102 | 0 | 0 | 13.5 | 1000 | 336.21 | 1000 | 380 | 63
4.2 | QQ Shoulder Helical MAR | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | true | false | S0.0 | Medium | 100 | true | 7.8 | 150 | 570 | 102 | 0 | 2 | 10.25 | 1000 | 267.11 | 1000
4.13 | GSIX Arm over Head Helical | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 100 |  |  |  |  | 102 | 0 | 0 | 14.49 | 1000 | 360.75 | 1000 | 270 | 63
4.7 | Hand nonHelical Routine | Series 2 Group 1 Scan Settings | 0.6 | Axial | SmartCollimation | true | false | S0.0 | Small Body | 100 | true | 7.8 | 100 | 450 | 102 | 0 | 0 | 9.92 | 1000 | 198.49 | 1000
4.14 | GSI Hand Helical SD3 | Series 2 Group 1 Scan Settings | 0.5 | Helical | 64 | false | false | S0.0 | Medium | 140 |  |  |  |  | 102 | 0 | 0 | 26.28 | 1000 | 619.14 | 1000 | 485 | 33
4.15 | Hand Tendon 140keV nonHelical naga | Series 2 Group 1 Scan Settings | 1 | Axial | SmartCollimation | true | false | S0.0 | Small Body | 140 |  |  |  |  | 102 | 0 | 0 | 16.6 | 1000 | 265.6 | 1000 | 200
4.16 | GSIX - Shoulder Routine | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 120 |  |  |  |  | 102 | 0 | 1 | 14.49 | 1000 | 360.75 | 1000 | 270 | 63
4.11 | Shunt CTA Routine SmartPrep | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.27 | 1000 | 648.33 | 1000
4.11 | Shunt CTA Routine SmartPrep | Series 3 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.32 | 1000 | 618.19 | 1000
4.11 | Shunt CTA Routine SmartPrep | Series 3 Group 2 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.27 | 1000 | 648.33 | 1000
4.17 | Hand Tendon 140keV Helical | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | true | false | S0.0 | Small Body | 140 | true | 3 | 200 | 405 | 102 | 0 | 0 | 16.46 | 1000 | 560.53 | 1000
5.13 | PA PV (29s 90s) Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 7.5 | 1000 | 278.23 | 1000
5.13 | PA PV (29s 90s) Routine | Series 2 Group 2 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 7.5 | 1000 | 199.52 | 1000
5.17 | ASO SmartPrep | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.27 | 1000 | 1244.09 | 1000
5.17 | ASO SmartPrep | Series 3 Group 1 Scan Settings | 0.8 | Helical | 64 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.18 | 1000 | 1199.88 | 1000
5.18 | AKA Go&Return SmartPrep | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.27 | 1000 | 648.33 | 1000
5.18 | AKA Go&Return SmartPrep | Series 3 Group 1 Scan Settings | 0.35 | Helical | 64 | false | false | S0.0 | Medium | 100 | true | 9.8 | 200 | 700 | 102 | 0 | 2 | 10.39 | 1000 | 634.35 | 1000
5.18 | AKA Go&Return SmartPrep | Series 3 Group 2 Scan Settings | 0.35 | Helical | 64 | false | false | S0.0 | Medium | 100 | true | 9.8 | 200 | 700 | 102 | 0 | 2 | 10.39 | 1000 | 634.35 | 1000
5.39 | CCTA + Function | Series 2 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Large | 120 | true | 35 | 100 | 200 | 102 | 0 | 1 | 2.66 | 1000 | 42.6 | 1000
5.39 | CCTA + Function | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | true | false | S0.0 | Cardiac Medium | 120 | true | 25 | 100 | 475 | 102 | 0 | 1 | 7.49 | 1000 | 104.91 | 1000
5.32 | GSIX Chest Low CA | Series 2 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 12.2 | 100 | 700 | 102 | 0 | 2 | 10.46 | 1000 | 613.5 | 1000 | 185 | 127
5.32 | GSIX Chest Low CA | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 317.43 | 1000
5.1 | QQ Chest Routine | Series 2 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 14.7 | 100 | 700 | 102 | 0 | 2 | 10.45 | 1000 | 612.89 | 1000
5.1 | QQ Chest Routine | Series 3 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 14.7 | 100 | 700 | 102 | 0 | 2 | 10.45 | 1000 | 403.84 | 1000
5.11 | PA + DV Pause | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.13 | 1000 | 376.12 | 1000
5.11 | PA + DV Pause | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 80 | true | 11.6 | 150 | 600 | 102 | 0 | 2 | 4.65 | 1000 | 562.63 | 1000
5.4 | CABG | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 17.3 | 50 | 250 | 102 | 0 | 1 | 1.53 | 1000 | 55.15 | 1000
5.4 | CABG | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Large | 120 | true | 35 | 100 | 200 | 102 | 0 | 1 | 2.66 | 1000 | 42.6 | 1000
5.4 | CABG | Series 4 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | true | false | S0.0 | Cardiac Medium | 120 | true | 25 | 100 | 475 | 102 | 0 | 1 | 6.47 | 1000 | 237.7 | 1000
5.42 | CABG + Function | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 17.3 | 50 | 250 | 102 | 0 | 1 | 1.51 | 1000 | 54.51 | 1000
5.42 | CABG + Function | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Large | 120 | true | 35 | 100 | 200 | 102 | 0 | 1 | 2.97 | 1000 | 47.45 | 1000
5.42 | CABG + Function | Series 4 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | true | false | S0.0 | Cardiac Medium | 120 | true | 25 | 100 | 475 | 102 | 0 | 1 | 7.11 | 1000 | 261.34 | 1000
5.19 | Ablation | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 17.3 | 50 | 250 | 103 | 0 | 1 | 1.51 | 1000 | 54.51 | 1000
5.19 | Ablation | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Medium | 120 | true | 25 | 100 | 740 | 103 | 0 | 1 | 25.34 | 1000 | 931.21 | 1000
5.2 | TAVI (Helical+Axial+Helical) | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 50 | 560 | 102 | 0 | 1 | 1.37 | 1000 | 104.44 | 1000
5.2 | TAVI (Helical+Axial+Helical) | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Large | 120 | true | 35 | 50 | 200 | 102 | 0 | 1 | 2.91 | 1000 | 46.53 | 1000
5.2 | TAVI (Helical+Axial+Helical) | Series 4 Group 1 Scan Settings | 0.28 | Helical | 64 | false | false | S0.0 | Medium | 100 | true | 14.7 | 300 | 720 | 102 | 0 | 0 | 2.92 | 1000 | 55.19 | 1000
5.2 | TAVI (Helical+Axial+Helical) | Series 4 Group 2 Scan Settings | 0.28 | Cardiac | 256 | false | false | S0.0 | Cardiac Medium | 120 | true | 13.9 | 100 | 740 | 102 | 0 | 0 | 14.5 | 1000 | 231.97 | 1000
5.2 | TAVI (Helical+Axial+Helical) | Series 4 Group 3 Scan Settings | 0.28 | Helical | 64 | false | false | S0.0 | Medium | 100 | true | 14.7 | 150 | 720 | 102 | 0 | 0 | 5.44 | 1000 | 276.54 | 1000
5.21 | TAVI (Coronary+Aorta no link) | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 150 | 700 | 102 | 0 | 2 | 1.53 | 1000 | 100.95 | 1000
5.21 | TAVI (Coronary+Aorta no link) | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Large | 120 | true | 35 | 50 | 200 | 102 | 0 | 1 | 2.61 | 1000 | 41.7 | 1000
5.21 | TAVI (Coronary+Aorta no link) | Series 4 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | true | false | S0.0 | Cardiac Medium | 120 | true | 25 | 100 | 475 | 102 | 0 | 0 | 35.67 | 1000 | 570.77 | 1000
5.21 | TAVI (Coronary+Aorta no link) | Series 4 Group 2 Scan Settings | 0.28 | Helical | 64 | false | false | S0.0 | Medium | 100 | true | 11.6 | 150 | 720 | 102 | 0 | 0 | 18.13 | 1000 | 1161.62 | 1000
5.22 | GSI - Chest | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.42 | 1000 | 303.61 | 1000 | 215 | 127
5.22 | GSI - Chest | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.42 | 1000 | 303.61 | 1000 | 215 | 127
5.23 | GSI - Chest - Pelvis | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.42 | 1000 | 530.84 | 1000 | 215 | 127
5.23 | GSI - Chest - Pelvis | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.42 | 1000 | 530.84 | 1000 | 215 | 127
5.24 | GSI - PA + DV | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.27 | 1000 | 648.33 | 1000
5.24 | GSI - PA + DV | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 326.23 | 1000 | 185 | 127
5.24 | GSI - PA + DV | Series 4 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 80 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 1065.32 | 1000 | 185 | 127
5.25 | GSI - Aorta CTA SmartPrep | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.15 | 1000 | 640.86 | 1000
5.25 | GSI - Aorta CTA SmartPrep | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.42 | 1000 | 530.84 | 1000 | 215 | 127
5.25 | GSI - Aorta CTA SmartPrep | Series 3 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.42 | 1000 | 530.84 | 1000 | 215 | 127
5.26 | GSI - ASO SmartPrep | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.15 | 1000 | 1229.74 | 1000
5.26 | GSI - ASO SmartPrep | Series 3 Group 1 Scan Settings | 0.8 | Helical | 64 | false | false | S0.0 | Medium | 120 |  |  |  |  | 102 | 0 | 2 | 10.5 | 1000 | 1238.47 | 1000 | 380 | 63
5.29 | GSIX PE+ DVT Routine | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 6.35 | 1000 | 244.94 | 1000 | 185 | 176
5.29 | GSIX PE+ DVT Routine | Series 2 Group 2 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 80 |  |  |  |  | 102 | 0 | 2 | 14.49 | 1000 | 1128.53 | 1000 | 270 | 63
5.44 | CARTO R to R | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 17.3 | 50 | 250 | 103 | 0 | 1 | 1.53 | 1000 | 55.15 | 1000
5.44 | CARTO R to R | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Medium | 120 | true | 25 | 100 | 740 | 103 | 0 | 1 | 34.82 | 1000 | 487.49 | 1000
5.4 | QQ Chest - Pelvis Routine | Series 2 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.44 | 1000 | 638.05 | 1000
5.4 | QQ Chest - Pelvis Routine | Series 3 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.44 | 1000 | 638.05 | 1000
5.2 | Chest (Low Dose) | Series 2 Group 1 Scan Settings | 0.35 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 30 | 50 | 150 | 102 | 0 | 2 | 1.75 | 1000 | 67.72 | 1000
5.5 | Chest - Pelvis (90sec 300sec) | Series 2 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 14.7 | 100 | 700 | 102 | 0 | 2 | 10.45 | 1000 | 665.15 | 1000
5.5 | Chest - Pelvis (90sec 300sec) | Series 3 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 14.7 | 100 | 700 | 102 | 0 | 2 | 10.45 | 1000 | 665.15 | 1000
5.5 | Chest - Pelvis (90sec300sec) | Series 4 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 14.7 | 100 | 700 | 102 | 0 | 2 | 10.45 | 1000 | 665.15 | 1000
5.1 | Chest - Pelvis Dynamic (40sec70sec180sec)Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.27 | 1000 | 268.29 | 1000
5.1 | Chest - Pelvis Dynamic (40sec70sec180sec)Routine | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.27 | 1000 | 648.33 | 1000
5.1 | Chest - Pelvis Dynamic (40sec70sec180sec)Routine | Series 3 Group 2 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.27 | 1000 | 679.15 | 1000
5.1 | Chest - Pelvis Dynamic (40sec70sec180sec)Routine | Series 3 Group 3 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.27 | 1000 | 648.33 | 1000
5.12 | Chest - Pelvis(40sec80sec)Routine | Series 2 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 14.7 | 100 | 700 | 102 | 0 | 2 | 10.45 | 1000 | 555.4 | 1000
5.12 | Chest - Pelvis(40sec80sec)Routine | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.27 | 1000 | 519.94 | 1000
5.12 | Chest - Pelvis(40sec80sec)Routine | Series 3 Group 2 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.27 | 1000 | 745.91 | 1000
5.14 | Aorta CTA Routine SmartPrep | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.28 | 1000 | 648.98 | 1000
5.14 | Aorta CTA Routine SmartPrep | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.28 | 1000 | 648.98 | 1000
5.14 | Aorta CTA Routine SmartPrep | Series 3 Group 2 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.28 | 1000 | 648.98 | 1000
5.37 | CCTA | Series 2 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Large | 120 | true | 35 | 100 | 200 | 102 | 0 | 1 | 2.66 | 1000 | 42.6 | 1000
5.37 | CCTA | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | true | false | S0.0 | Cardiac Medium | 120 | true | 25 | 100 | 475 | 102 | 0 | 1 | 7.49 | 1000 | 104.91 | 1000
5.43 | CARTO Systole | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 17.3 | 50 | 250 | 103 | 0 | 1 | 1.53 | 1000 | 55.15 | 1000
5.43 | CARTO Systole | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Medium | 120 | true | 25 | 100 | 740 | 103 | 0 | 1 | 6.83 | 1000 | 95.59 | 1000
5.6 | QQ PE + DVT Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.13 | 1000 | 376.12 | 1000
5.6 | QQ PE + DVT Routine | Series 2 Group 2 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 80 | true | 11.6 | 150 | 600 | 102 | 0 | 2 | 5.47 | 1000 | 443.67 | 1000
5.16 | Chest HRCT | Series 1 Group 1 Scan Settings | 0.5 | Axial | 128 | true | false | S0.0 | Medium | 100 | true | 14.7 | 100 | 570 | 102 | 0 | 2 | 10.47 | 1000 | 83.78 | 1000
5.3 | QQ Chest Routine Body Assist | Series 2 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 14.7 | 100 | 720 | 102 | 0 | 2 | 10.46 | 1000 | 613.5 | 1000
5.3 | QQ Chest Routine Body Assist | Series 3 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 14.7 | 100 | 700 | 102 | 0 | 2 | 10.46 | 1000 | 404.25 | 1000
5.33 | GSIX Chest-Pelvis Low CA | Series 2 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 12.2 | 100 | 700 | 102 | 0 | 2 | 10.46 | 1000 | 613.5 | 1000
5.33 | GSIX Chest-Pelvis Low CA | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 317.43 | 1000 | 185 | 127
5.9 | P-Abd CE Chest-Pelvis Routine | Series 2 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 14.7 | 100 | 700 | 102 | 0 | 2 | 10.46 | 1000 | 665.81 | 1000
5.9 | P-Abd CE Chest-Pelvis Routine | Series 3 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 14.7 | 100 | 700 | 102 | 0 | 2 | 10.46 | 1000 | 665.81 | 1000
5.45 | GSIX - Aorta CTA SmartPrep | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.27 | 1000 | 576.43 | 1000
5.45 | GSIX - Aorta CTA SmartPrep | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
5.45 | GSIX - Aorta CTA SmartPrep | Series 3 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 21.21 | 1000 | 1687.48 | 1000 | 435 | 127
5.45 | GSIX - Aorta CTA SmartPrep | Series 3 Group 3 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
5.31 | GSIX PA PV (29s 90s) | Series 2 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 14.7 | 100 | 700 | 102 | 0 | 2 | 10.46 | 1000 | 467.02 | 1000
5.31 | GSIX PA PV (29s 90s) | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 979.21 | 1000 | 475 | 127
5.31 | GSIX PA PV (29s 90s) | Series 3 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 979.21 | 1000 | 475 | 127
5.35 | GSIX Aorta CTA(CTA 100kV Delay Dual) | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.28 | 1000 | 648.98 | 1000
5.35 | GSIX Aorta CTA(CTA 100kV Delay Dual) | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11 | 100 | 720 | 102 | 0 | 2 | 7.81 | 1000 | 492.88 | 1000
5.35 | GSIX Aorta CTA(CTA 100kV Delay Dual) | Series 3 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 555 | 1000 | 185 | 127
5.38 | CCTA + 4D | Series 2 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Large | 120 | true | 35 | 100 | 200 | 102 | 0 | 1 | 2.66 | 1000 | 42.6 | 1000
5.38 | CCTA + 4D | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | true | false | S0.0 | Cardiac Medium | 120 | true | 25 | 100 | 475 | 102 | 0 | 1 | 35.67 | 1000 | 499.42 | 1000
5.34 | GSIX PA+PV+DVT(7s14s) | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 5.58 | 1000 | 201.44 | 1000 | 195 | 127
5.34 | GSIX PA+PV+DVT(7s14s) | Series 2 Group 2 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 5.58 | 1000 | 201.44 | 1000 | 195 | 127
5.34 | GSIX PA+PV+DVT(7s14s) | Series 2 Group 3 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 80 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 713.37 | 1000 | 185 | 127
5.3 | GSIX PE+ DVT (20s 80s 180s) | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 103 | 0 | 2 | 5.58 | 1000 | 201.44 | 1000 | 195 | 127
5.3 | GSIX PE+ DVT (20s 80s 180s) | Series 2 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 103 | 0 | 2 | 8.8 | 1000 | 317.43 | 1000 | 185 | 127
5.3 | GSIX PE+ DVT (20s 80s 180s) | Series 2 Group 3 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 80 |  |  |  |  | 102 | 0 | 2 | 14.49 | 1000 | 1128.53 | 1000 | 270 | 63
5.46 | GSIX PE+ DVT(100kv) | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 103 | 0 | 2 | 5.58 | 1000 | 201.44 | 1000 | 195 | 127
5.46 | GSIX PE+ DVT(100kv) | Series 2 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 150 | 720 | 102 | 0 | 2 | 7.8 | 1000 | 632.8 | 1000
5.8 | P+CE Chest-Pelvis Routine | Series 2 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.44 | 1000 | 638.05 | 1000
5.8 | P+CE Chest-Pelvis Routine | Series 3 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.44 | 1000 | 638.05 | 1000
5.28 | GSIX Chest-Pelvis CEonly | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 317.43 | 1000 | 185 | 127
5.15 | Aorta CTA (50s 90s) SmartPrep | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.28 | 1000 | 648.98 | 1000
5.15 | Aorta CTA (50s 90s) SmartPrep | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.28 | 1000 | 648.98 | 1000
5.15 | Aorta CTA (50s 90s) SmartPrep | Series 3 Group 2 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.28 | 1000 | 648.98 | 1000
5.15 | Aorta CTA (50s 90s) SmartPrep | Series 3 Group 3 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.28 | 1000 | 648.98 | 1000
5.27 | GSIX Chest-Pelvi P+CE | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.05 | 1000 | 1325.64 | 1000
5.27 | GSIX Chest-Pelvi P+CE | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 493.41 | 1000 | 185 | 127
5.47 | CABG Special | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 14 | 100 | 700 | 102 | 0 | 1 | 1.53 | 1000 | 50.62 | 1000
5.47 | CABG Special | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Large | 120 | true | 35 | 100 | 200 | 102 | 0 | 1 | 2.66 | 1000 | 42.6 | 1000
5.47 | CABG Special | Series 4 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | true | false | S0.0 | Cardiac Medium | 120 | true | 25 | 100 | 470 | 102 | 0 | 1 | 8.56 | 1000 | 209.8 | 1000
5.47 | CABG Special | Series 5 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 10.87 | 1000 | 555.46 | 1000 | 230 | 127
5.5 | Trauma Full Body DMPR | Series 2 Group 1 Scan Settings | 1 | Axial | SmartCollimation | false | false | S0.0 | Adult Head | 120 |  |  |  |  | 101 | 0 | 0 | 56.45 | 1000 | 903.13 | 1000 | 340
5.5 | Trauma Full Body DMPR | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | true | false | S0.0 | Large Body | 120 | true | 8.5 | 80 | 430 | 101 | 0 | 0 | 17.28 | 1000 | 547.84 | 1000
5.5 | Trauma Full Body DMPR | Series 4 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 16.6 | 100 | 720 | 101 | 0 | 0 | 5.28 | 1000 | 187.93 | 1000
5.5 | Trauma Full Body DMPR | Series 4 Group 2 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Large Body | 120 | true | 12.3 | 100 | 740 | 101 | 0 | 0 | 17.01 | 1000 | 776.09 | 1000
5.36 | GSIX Aorta CTA(CTA 50s 90s)90s only Dual | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.28 | 1000 | 648.98 | 1000
5.36 | GSIX Aorta CTA(CTA 50s 90s)90s only Dual | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11 | 100 | 720 | 102 | 0 | 2 | 7.81 | 1000 | 492.88 | 1000
5.36 | GSIX Aorta CTA(CTA 50s 90s)90s only Dual | Series 3 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11 | 100 | 720 | 102 | 0 | 2 | 7.81 | 1000 | 492.88 | 1000
5.36 | GSIX Aorta CTA(CTA 50s 90s)90s only Dual | Series 3 Group 3 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 555 | 1000 | 185 | 127
5.48 | CABG Special2 | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 14 | 100 | 700 | 102 | 0 | 1 | 1.53 | 1000 | 50.62 | 1000
5.48 | CABG Special2 | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Large | 120 | true | 35 | 100 | 200 | 102 | 0 | 1 | 2.66 | 1000 | 42.6 | 1000
5.48 | CABG Special2 | Series 4 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | true | false | S0.0 | Cardiac Medium | 120 | true | 25 | 100 | 470 | 102 | 0 | 1 | 8.56 | 1000 | 209.8 | 1000
5.48 | CABG Special2 | Series 5 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 10.87 | 1000 | 555.46 | 1000 | 230 | 127
5.48 | CABG Special2 | Series 5 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 10.87 | 1000 | 555.46 | 1000 | 230 | 127
5.51 | GSIX Chest-Pelvi CE all dual | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 804.67 | 1000 | 475 | 127
5.51 | GSIX Chest-Pelvi CE all dual | Series 2 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1270.1 | 1000 | 475 | 127
5.51 | GSIX Chest-Pelvi CE all dual | Series 2 Group 3 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 804.67 | 1000 | 475 | 127
5.7 | QQ Full FOV Chest - Pelvis 80cm | Series 2 Group 1 Scan Settings | 0.5 | Helical | 64 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 9.52 | 1000 | 564.71 | 1000
5.7 | QQ Full FOV Chest - Pelvis 80cm | Series 3 Group 1 Scan Settings | 0.5 | Helical | 64 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 9.52 | 1000 | 564.71 | 1000
5.41 | CABG+CTA+180sDE | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 14 | 100 | 700 | 102 | 0 | 1 | 1.53 | 1000 | 109.46 | 1000
5.41 | CABG+CTA+180sDE | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Large | 120 | true | 35 | 100 | 200 | 102 | 0 | 1 | 2.66 | 1000 | 42.6 | 1000
5.41 | CABG+CTA+180sDE | Series 4 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | true | false | S0.0 | Cardiac Medium | 120 | true | 25 | 100 | 450 | 102 | 0 | 1 | 8.49 | 1000 | 237.58 | 1000
5.41 | CABG+CTA+180sDE | Series 4 Group 2 Scan Settings | 0.8 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 14 | 100 | 720 | 102 | 0 | 1 | 5 | 1000 | 360.74 | 1000
5.41 | CABG+CTA+180sDE | Series 5 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 10.87 | 1000 | 783.83 | 1000 | 230 | 127
5.49 | GSIX - Aorta CTA SmartPrep (120sDE) | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.28 | 1000 | 577.01 | 1000
5.49 | GSIX - Aorta CTA SmartPrep (120sDE) | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
5.49 | GSIX - Aorta CTA SmartPrep (120sDE) | Series 3 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 21.21 | 1000 | 1687.48 | 1000 | 435 | 127
5.52 | GSIX Chest-pelvi CTA (All Dual) | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 635 | 102 | 0 | 2 | 15.62 | 1000 | 1891.56 | 1000
5.52 | GSIX Chest-pelvi CTA (All Dual) | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 1065.32 | 1000 | 185 | 127
5.52 | GSIX Chest-pelvi CTA (All Dual) | Series 4 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 1065.32 | 1000 | 185 | 127
6.11 | Abd-FAT Routine | Series 1 Group 1 Scan Settings | 0.35 | Axial | 8 | false | false | S0.0 | Medium | 100 |  |  |  |  | 102 | 0 | 2 | 2.42 | 1000 | 1.21 | 1000 | 100
6.18 | GSIX Abdomen-Pelvi P+CE | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.05 | 1000 | 1325.64 | 1000
6.18 | GSIX Abdomen-Pelvi P+CE | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 493.41 | 1000 | 185 | 127
6.14 | GSI - Abdomen-Pelvis | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 449.41 | 1000 | 215 | 127
6.14 | GSI - Abdomen-Pelvis | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 493.41 | 1000 | 215 | 127
6.1 | QQ Abdomen-Pelvis Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.04 | 1000 | 1024.62 | 1000
6.1 | QQ Abdomen-Pelvis Routine | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.04 | 1000 | 1124.84 | 1000
6.9 | Kunishima 4Phase (Smart Prep) Expiration | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 103 | 0 | 2 | 20.02 | 1000 | 1023.6 | 1000
6.9 | Kunishima 4Phase (Smart Prep) Expiration | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 103 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.9 | Kunishima 4Phase (Smart Prep) Expiration | Series 3 Group 2 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 103 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.9 | Kunishima 4Phase (Smart Prep) Expiration | Series 3 Group 3 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 103 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.9 | Kunishima 4Phase (Smart Prep) Expiration | Series 3 Group 4 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 103 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.7 | RCC 3Phase (30sec80sec300sec)PauseRoutine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1023.6 | 1000
6.7 | RCC 3Phase (30sec80sec300sec)PauseRoutine | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.7 | RCC 3Phase (30sec80sec300sec)PauseRoutine | Series 3 Group 2 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.7 | RCC 3Phase (30sec80sec300sec)PauseRoutine | Series 4 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.6 | RCC 3Phase (30sec80sec300sec)Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1023.6 | 1000
6.6 | RCC 3Phase (30sec80sec300sec)Routine | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.6 | RCC 3Phase (30sec80sec300sec)Routine | Series 3 Group 2 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.6 | RCC 3Phase (30sec80sec300sec)Routine | Series 3 Group 3 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.5 | HCC 4Phase (25sec40sec70sec180sec) | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1023.6 | 1000
6.5 | HCC 4Phase (25sec40sec70sec180sec) | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.5 | HCC 4Phase (25sec40sec70sec180sec) | Series 3 Group 2 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.5 | HCC 4Phase (25sec40sec70sec180sec) | Series 3 Group 3 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.5 | HCC 4Phase (25sec40sec70sec180sec) | Series 3 Group 4 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.4 | HCC 3Phase (40sec70sec180sec)Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1023.6 | 1000
6.4 | HCC 3Phase (40sec70sec180sec)Routine | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.4 | HCC 3Phase (40sec70sec180sec)Routine | Series 3 Group 2 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.4 | HCC 3Phase (40sec70sec180sec)Routine | Series 3 Group 3 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.12 | Abdomen 3Phase SmartPrep | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1023.6 | 1000
6.12 | Abdomen 3Phase SmartPrep | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.12 | Abdomen 3Phase SmartPrep | Series 3 Group 2 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.12 | Abdomen 3Phase SmartPrep | Series 3 Group 3 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.13 | Abdomen 4Phase SmartPrep | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1023.6 | 1000
6.13 | Abdomen 4Phase SmartPrep | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.13 | Abdomen 4Phase SmartPrep | Series 3 Group 2 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.13 | Abdomen 4Phase SmartPrep | Series 3 Group 3 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.13 | Abdomen 4Phase SmartPrep | Series 3 Group 4 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.15 | GSI - Abdomen 2Phase SmartPrep | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 449.41 | 1000 | 215 | 127
6.15 | GSI - Abdomen 2Phase SmartPrep | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 493.41 | 1000 | 215 | 127
6.15 | GSI - Abdomen 2Phase SmartPrep | Series 3 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 493.41 | 1000 | 215 | 127
6.16 | GSI - Abdomen 3Phase SmartPrep | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 449.41 | 1000 | 185 | 127
6.16 | GSI - Abdomen 3Phase SmartPrep | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 493.41 | 1000 | 185 | 127
6.16 | GSI - Abdomen 3Phase SmartPrep | Series 3 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 493.41 | 1000 | 185 | 127
6.16 | GSI - Abdomen 3Phase SmartPrep | Series 3 Group 3 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 493.41 | 1000 | 185 | 127
6.17 | GSI - Abdomen 4Phase SmartPrep | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 449.41 | 1000 | 185 | 127
6.17 | GSI - Abdomen 4Phase SmartPrep | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 493.41 | 1000 | 185 | 127
6.17 | GSI - Abdomen 4Phase SmartPrep | Series 3 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 493.41 | 1000 | 185 | 127
6.17 | GSI - Abdomen 4Phase SmartPrep | Series 3 Group 3 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 493.41 | 1000 | 185 | 127
6.17 | GSI - Abdomen 4Phase SmartPrep | Series 3 Group 4 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 493.41 | 1000 | 185 | 127
6.36 | RT Routin | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Medium | 120 | false | 9 | 150 | 515 | 102 | 0 | 0 | 38.98 | 1000 | 1821.35 | 1000
6.36 | RT Routin | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Medium | 120 | false | 9 | 150 | 515 | 102 | 0 | 0 | 38.98 | 1000 | 1821.35 | 1000
6.23 | GSIX Abdomen-Pelvis Low CA | Series 2 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 12.2 | 100 | 700 | 102 | 0 | 2 | 10.45 | 1000 | 612.89 | 1000
6.23 | GSIX Abdomen-Pelvis Low CA | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 317.43 | 1000 | 185 | 127
6.24 | GSIX Abdomen-Pelvi Liver Dynamic | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.03 | 1000 | 823.59 | 1000
6.24 | GSIX Abdomen-Pelvi Liver Dynamic | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
6.24 | GSIX Abdomen-Pelvi Liver Dynamic | Series 3 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
6.24 | GSIX Abdomen-Pelvi Liver Dynamic | Series 3 Group 3 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
6.21 | GSIX Dynamic(40s 70s 180s) 70s onlyDE | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 9.6 | 100 | 635 | 102 | 0 | 2 | 20.05 | 1000 | 1125.15 | 1000
6.21 | GSIX Dynamic(40s 70s 180s) 70s onlyDE | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 635 | 102 | 0 | 2 | 20.05 | 1000 | 1125.15 | 1000
6.21 | GSIX Dynamic(40s 70s 180s) 70s onlyDE | Series 3 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
6.21 | GSIX Dynamic(40s 70s 180s) 70s onlyDE | Series 3 Group 3 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 635 | 102 | 0 | 2 | 20.05 | 1000 | 1125.15 | 1000
6.28 | GSIX Pelvi umeda protocol CTA(40s 70s 180s) 70s onlyDE | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
6.28 | GSIX Pelvi umeda protocol CTA(40s 70s 180s) 70s onlyDE | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 635 | 102 | 0 | 2 | 20.05 | 1000 | 1125.15 | 1000
6.28 | GSIX Pelvi umeda protocol CTA(40s 70s 180s) 70s onlyDE | Series 3 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 635 | 102 | 0 | 2 | 20.05 | 1000 | 1125.15 | 1000
6.28 | GSIX Pelvi umeda protocol CTA(40s 70s 180s) 70s onlyDE | Series 3 Group 3 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
6.28 | GSIX Pelvi umeda protocol CTA(40s 70s 180s) 70s onlyDE | Series 3 Group 4 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 635 | 102 | 0 | 2 | 20.05 | 1000 | 1125.15 | 1000
6.31 | GSIX Pelvi umeda protocol CTA (22s 50s 90s) 90s onlyDE | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
6.31 | GSIX Pelvi umeda protocol CTA (22s 50s 90s) 90s onlyDE | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 635 | 102 | 0 | 2 | 20.05 | 1000 | 1125.15 | 1000
6.31 | GSIX Pelvi umeda protocol CTA (22s 50s 90s) 90s onlyDE | Series 3 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 635 | 102 | 0 | 2 | 20.05 | 1000 | 1125.15 | 1000
6.31 | GSIX Pelvi umeda protocol CTA (22s 50s 90s) 90s onlyDE | Series 3 Group 3 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
6.33 | RT HCC 4D 2Phase 40s180s | Series 2 Group 1 Scan Settings | 1 | Cine | 256 | false | false | S0.0 | Medium | 80 | true | 11.6 | 100 | 600 | 102 | 0 | 2 | 52.12 | 1000 | 833.85 | 1000
6.33 | RT HCC 4D 2Phase 40s180s | Series 2 Group 2 Scan Settings | 0.5 | Cine | 256 | false | false | S0.0 | Medium | 80 | true | 11.6 | 100 | 600 | 102 | 0 | 2 | 64.38 | 1000 | 1030.09 | 1000
6.25 | GSI - Abdomen 3Phase SmartPrep - Copy | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 449.41 | 1000 | 185 | 127
6.25 | GSI - Abdomen 3Phase SmartPrep - Copy | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 493.41 | 1000 | 185 | 127
6.25 | GSI - Abdomen 3Phase SmartPrep - Copy | Series 3 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 493.41 | 1000 | 185 | 127
6.25 | GSI - Abdomen 3Phase SmartPrep - Copy | Series 3 Group 3 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 493.41 | 1000 | 185 | 127
6.3 | GSIX Pelvi umeda protocol CTA (50s 90s) 90s onlyDE | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
6.3 | GSIX Pelvi umeda protocol CTA (50s 90s) 90s onlyDE | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 635 | 102 | 0 | 2 | 20.05 | 1000 | 1125.15 | 1000
6.3 | GSIX Pelvi umeda protocol CTA (50s 90s) 90s onlyDE | Series 3 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 635 | 102 | 0 | 2 | 20.05 | 1000 | 1125.15 | 1000
6.3 | GSIX Pelvi umeda protocol CTA (50s 90s) 90s onlyDE | Series 3 Group 3 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
6.3 | Abdomen 2Phase (40sec70sec) | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1023.6 | 1000
6.3 | Abdomen 2Phase (40sec70sec) | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.3 | Abdomen 2Phase (40sec70sec) | Series 3 Group 2 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.2 | GSIX Chest-Pelvi - CE only | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
6.8 | Abd Stone 2Phase (90sec5min)Pause | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1023.6 | 1000
6.8 | Abd Stone 2Phase (90sec5min)Pause | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.8 | Abd Stone 2Phase (90sec5min)Pause | Series 4 Group 1 Scan Settings | 0.5 | Axial | 8 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 101 | 0 | 2 | 3.15 | 1000 | 1.57 | 1000 | 100
6.8 | Abd Stone 2Phase (90sec5min)Pause | Series 5 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 1123.72 | 1000
6.1 | Abd Aorta CTA Routine SmartPrep | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.27 | 1000 | 648.33 | 1000
6.1 | Abd Aorta CTA Routine SmartPrep | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.27 | 1000 | 648.33 | 1000
6.1 | Abd Aorta CTA Routine SmartPrep | Series 3 Group 2 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.27 | 1000 | 648.33 | 1000
6.19 | GSIX Abdomen-Pelvi - CE only | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
6.34 | RT Radiation Routine | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Large Body | 120 | true | 14 | 100 | 720 | 102 | 0 | 2 | 25.58 | 1000 | 1226.1 | 1000
6.34 | RT Radiation Routine | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.04 | 1000 | 1124.84 | 1000
6.22 | GSIX Radiation Planning CT | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 100 |  |  |  |  | 102 | 0 | 2 | 26.01 | 1000 | 1375.81 | 1000 | 475 | 63
6.35 | RT HCC 4D Test | Series 2 Group 1 Scan Settings | 0.28 | Cine | 256 | false | false | S0.0 | Medium | 120 | true | 16.4 | 100 | 600 | 102 | 0 | 2 | 137.15 | 1000 | 2194.37 | 1000
6.35 | RT HCC 4D Test | Series 3 Group 1 Scan Settings | 0.28 | Axial | 256 | false | false | S0.0 | Medium | 120 | true | 16.4 | 100 | 600 | 102 | 0 | 2 | 140.76 | 1000 | 2252.14 | 1000
6.26 | GSIX Abdomen-Pelvi Liver Dynamic all dual | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 804.67 | 1000 | 475 | 127
6.26 | GSIX Abdomen-Pelvi Liver Dynamic all dual | Series 2 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1270.1 | 1000 | 475 | 127
6.26 | GSIX Abdomen-Pelvi Liver Dynamic all dual | Series 2 Group 3 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 804.67 | 1000 | 475 | 127
6.2 | Abdomen-Pelvis Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.04 | 1000 | 1024.62 | 1000
6.2 | Abdomen-Pelvis Routine | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.04 | 1000 | 1124.84 | 1000
6.27 | Abdomen-Pelvis 4D | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 720 | 102 | 0 | 2 | 20.04 | 1000 | 1024.62 | 1000
6.27 | Abdomen-Pelvis 4D | Series 3 Group 1 Scan Settings | 0.28 | Axial | 256 | false | false | S0.0 | Medium | 100 | true | 15.3 | 100 | 600 | 102 | 0 | 2 | 254.31 | 1000 | 4068.98 | 1000
6.27 | Abdomen-Pelvis 4D | Series 3 Group 2 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 600 | 102 | 0 | 2 | 6.39 | 1000 | 294.77 | 1000
6.32 | GSIX Abd CTA (All Dual) | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 635 | 102 | 0 | 2 | 15.62 | 1000 | 1891.56 | 1000
6.32 | GSIX Abd CTA (All Dual) | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | 185 | 127 | 102 | 0 | 2 | 8.8 | 1000 | 1065.32 | 1000
6.32 | GSIX Abd CTA (All Dual) | Series 4 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | 185 | 127 | 102 | 0 | 2 | 8.8 | 1000 | 1065.32 | 1000
6.29 | GSIX Pelvi umeda protocol CTA(40s 70s 180s) 40s 70s DE | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
6.29 | GSIX Pelvi umeda protocol CTA(40s 70s 180s) 40s 70s DE | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
6.29 | GSIX Pelvi umeda protocol CTA(40s 70s 180s) 40s 70s DE | Series 3 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
6.29 | GSIX Pelvi umeda protocol CTA(40s 70s 180s) 40s 70s DE | Series 3 Group 3 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.4 | 100 | 635 | 102 | 0 | 2 | 20.05 | 1000 | 1125.15 | 1000
6.37 | RT HCC 2Phase 40s180s | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Medium | 120 | true | 9 | 100 | 600 | 102 | 0 | 2 | 4.35 | 1000 | 157.2 | 1000
6.37 | RT HCC 2Phase 40s180s | Series 2 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Medium | 120 | true | 9 | 100 | 600 | 102 | 0 | 2 | 8.34 | 1000 | 301.31 | 1000
6.39 | GSIX Pelvi CTA(40s 70s 180s) all DE | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
6.39 | GSIX Pelvi CTA(40s 70s 180s) all DE | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
6.39 | GSIX Pelvi CTA(40s 70s 180s) all DE | Series 3 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
6.39 | GSIX Pelvi CTA(40s 70s 180s) all DE | Series 3 Group 3 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1305.01 | 1000 | 475 | 127
7.2 | QQ Th-Spine Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Small Body | 100 | true | 11.6 | 150 | 700 | 102 | 0 | 2 | 8.66 | 1000 | 295.1 | 1000
7.4 | QQ Whole Spine Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Small Body | 100 | true | 11.6 | 150 | 700 | 102 | 0 | 2 | 8.66 | 1000 | 416.36 | 1000
7.5 | C-Spine non Helical Routine | Series 2 Group 1 Scan Settings | 1 | Axial | SmartCollimation | false | false | S0.0 | Small Body | 100 | true | 5 | 150 | 700 | 101 | 0 | 0 | 9.36 | 1000 | 262.05 | 1000
7.1 | QQ C-Spine Helical MAR Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Small Body | 100 | true | 5 | 150 | 700 | 102 | 0 | 2 | 8.66 | 1000 | 260.45 | 1000
7.3 | QQ L-Spine Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Small Body | 100 | true | 11.6 | 150 | 700 | 102 | 0 | 2 | 8.66 | 1000 | 295.1 | 1000
7.1 | GSI - Th-Spine | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 120 |  |  |  |  | 101 | 0 | 0 | 14.49 | 1000 | 418.69 | 1000 | 380 | 63
7.11 | GSI - L-Spine | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 120 |  |  |  |  | 101 | 0 | 0 | 14.49 | 1000 | 534.59 | 1000 | 380 | 63
7.12 | GSI - Whole Spine | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 120 |  |  |  |  | 101 | 0 | 0 | 14.49 | 1000 | 679.45 | 1000 | 380 | 63
7.7 | GSIX Th-Spine Routine | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 100 |  |  |  |  | 102 | 0 | 2 | 14.49 | 1000 | 476.64 | 1000 | 270 | 63
7.6 | GSIX L-Spine Routine | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 100 |  |  |  |  | 102 | 0 | 2 | 14.49 | 1000 | 476.64 | 1000 | 270 | 63
7.8 | GSIX C-Spine Routine | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 100 |  |  |  |  | 102 | 0 | 2 | 14.49 | 1000 | 476.64 | 1000 | 270 | 63
7.9 | GSIX Whole-Spine Routine | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 100 |  |  |  |  | 102 | 0 | 2 | 14.49 | 1000 | 476.64 | 1000 | 270 | 63
8.4 | QQ Hip Joint Routine MAR | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 150 | 720 | 102 | 0 | 2 | 20.02 | 1000 | 923.48 | 1000
8.2 | GSI - Pelvis | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 120 |  |  |  |  | 102 | 0 | 2 | 13.5 | 1000 | 444.22 | 1000 | 380 | 63
8.2 | GSI - Pelvis | Series 3 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 120 |  |  |  |  | 102 | 0 | 2 | 13.5 | 1000 | 444.22 | 1000 | 380 | 63
8.3 | GSI - Hip Joint | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 120 |  |  |  |  | 101 | 0 | 0 | 13.5 | 1000 | 390.21 | 1000 | 380 | 63
8.1 | QQ Pelvis Rotine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 150 | 720 | 102 | 0 | 2 | 20.04 | 1000 | 924.41 | 1000
8.1 | QQ Pelvis Rotine | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 150 | 720 | 102 | 0 | 2 | 20.04 | 1000 | 924.41 | 1000
8.5 | GSIX Hip Joint Routine MAR | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 23.27 | 1000 | 1072.29 | 1000 | 475 | 127
9.1 | QQ Lower Extremity Heiical Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Medium | 100 | true | 4.1 | 100 | 570 | 101 | 0 | 0 | 10.18 | 1000 | 538.5 | 1000
9.15 | nonHelical 8slice Routine | Series 2 Group 1 Scan Settings | 1 | Axial | 8 | false | false | S0.0 | Medium | 100 |  |  |  |  | 102 | 0 | 0 | 5.53 | 1000 | 2.77 | 1000 | 80
9.15 | nonHelical 8slice Routine | Series 2 Group 2 Scan Settings | 1 | Axial | 8 | false | false | S0.0 | Medium | 100 |  |  |  |  | 102 | 0 | 0 | 5.53 | 1000 | 2.77 | 1000 | 80
9.15 | nonHelical 8slice Routine | Series 2 Group 3 Scan Settings | 1 | Axial | 8 | false | false | S0.0 | Medium | 100 |  |  |  |  | 102 | 0 | 0 | 5.53 | 1000 | 2.77 | 1000 | 80
9.15 | nonHelical 8slice Routine | Series 2 Group 4 Scan Settings | 1 | Axial | 8 | false | false | S0.0 | Medium | 100 |  |  |  |  | 102 | 0 | 0 | 5.53 | 1000 | 2.77 | 1000 | 80
9.15 | nonHelical 8slice Routine | Series 2 Group 5 Scan Settings | 1 | Axial | 8 | false | false | S0.0 | Medium | 100 |  |  |  |  | 102 | 0 | 0 | 5.53 | 1000 | 2.77 | 1000 | 80
9.15 | nonHelical 8slice Routine | Series 2 Group 6 Scan Settings | 1 | Axial | 8 | false | false | S0.0 | Medium | 100 |  |  |  |  | 102 | 0 | 0 | 5.53 | 1000 | 2.77 | 1000 | 80
9.15 | nonHelical 8slice Routine | Series 2 Group 7 Scan Settings | 1 | Axial | 8 | false | false | S0.0 | Medium | 100 |  |  |  |  | 102 | 0 | 0 | 5.53 | 1000 | 2.77 | 1000 | 80
9.15 | nonHelical 8slice Routine | Series 2 Group 8 Scan Settings | 1 | Axial | 8 | false | false | S0.0 | Medium | 100 |  |  |  |  | 102 | 0 | 0 | 5.53 | 1000 | 2.77 | 1000 | 80
9.2 | QQ Knee/Ankle Helical HR Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | true | false | S0.0 | Medium | 100 | true | 4.1 | 100 | 450 | 101 | 0 | 0 | 10.25 | 1000 | 267.11 | 1000
9.3 | Knee/Ankle non-Herical HR Routine | Series 2 Group 1 Scan Settings | 1 | Axial | SmartCollimation | true | false | S0.0 | Medium | 100 | true | 4.1 | 100 | 500 | 102 | 0 | 0 | 12.15 | 1000 | 291.65 | 1000
9.16 | ASO SmartPrep(CTA 90s 180s) | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.27 | 1000 | 1434.11 | 1000
9.16 | ASO SmartPrep(CTA 90s 180s) | Series 3 Group 1 Scan Settings | 0.8 | Helical | 64 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.18 | 1000 | 1388.12 | 1000
9.16 | ASO SmartPrep(CTA 90s 180s) | Series 3 Group 2 Scan Settings | 0.8 | Helical | 64 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.18 | 1000 | 1388.12 | 1000
9.16 | ASO SmartPrep(CTA 90s 180s) | Series 3 Group 3 Scan Settings | 0.8 | Helical | 64 | false | false | S0.0 | Medium | 100 | true | 11.6 | 100 | 700 | 102 | 0 | 2 | 10.18 | 1000 | 1388.12 | 1000
9.13 | Ankle Axial | Series 2 Group 1 Scan Settings | 1 | Axial | SmartCollimation | true | false | S0.0 | Medium | 100 | true | 4.1 | 100 | 500 | 102 | 0 | 0 | 12.15 | 1000 | 291.65 | 1000
9.14 | Ankle Helical | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | true | false | S0.0 | Medium | 100 | true | 4.1 | 100 | 450 | 101 | 0 | 0 | 10.25 | 1000 | 267.11 | 1000
9.9 | GSI - Lower Extremity | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 101 | 0 | 0 | 8.42 | 1000 | 505.59 | 1000 | 215 | 127
9.1 | GSI - Knee | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 120 |  |  |  |  | 101 | 0 | 0 | 13.5 | 1000 | 309.21 | 1000 | 380 | 63
9.5 | GSIX ASO Prep(CTA+Return all dual) | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 635 | 102 | 0 | 2 | 15.6 | 1000 | 1889.67 | 1000
9.5 | GSIX ASO Prep(CTA+Return all dual) | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 1065.32 | 1000 | 185 | 127
9.5 | GSIX ASO Prep(CTA+Return all dual) | Series 3 Group 2 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 1065.32 | 1000 | 185 | 127
9.5 | GSIX ASO Prep(CTA+Return all dual) | Series 4 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 1065.32 | 1000 | 185 | 127
9.8 | ASO 80keV SmartPrep Routine | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 80 | true | 11.6 | 100 | 600 | 102 | 0 | 2 | 3.98 | 1000 | 481.93 | 1000
9.8 | ASO 80keV SmartPrep Routine | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 80 | true | 11.6 | 100 | 600 | 102 | 0 | 2 | 3.98 | 1000 | 481.93 | 1000
9.8 | ASO 80keV SmartPrep Routine | Series 4 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 80 | true | 11.6 | 100 | 600 | 102 | 0 | 2 | 3.98 | 1000 | 481.93 | 1000
9.4 | GSIX ASO Prep (CE All Dual) | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 635 | 102 | 0 | 2 | 15.6 | 1000 | 1889.67 | 1000
9.4 | GSIX ASO Prep (CE All Dual) | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 1065.32 | 1000 | 185 | 127
9.4 | GSIX ASO Prep (CE All Dual) | Series 4 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 2 | 8.8 | 1000 | 1065.32 | 1000 | 185 | 127
9.11 | GSI - Ankle | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 120 |  |  |  |  | 101 | 0 | 0 | 13.5 | 1000 | 309.21 | 1000 | 380 | 63
9.12 | GSIX - Knee Routine | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Medium | 120 |  |  |  |  | 101 | 0 | 0 | 23.71 | 1000 | 697.15 | 1000 | 435 | 63
9.7 | ASO SmartPrep Routine | Series 2 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 635 | 102 | 0 | 2 | 7.8 | 1000 | 944.83 | 1000
9.7 | ASO SmartPrep Routine | Series 3 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 600 | 102 | 0 | 2 | 7.8 | 1000 | 944.83 | 1000
9.7 | ASO SmartPrep Routine | Series 4 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 100 | true | 11.6 | 100 | 600 | 102 | 0 | 2 | 7.8 | 1000 | 944.83 | 1000
9.6 | nonHelical 8slice Routine - 80cm | Series 2 Group 1 Scan Settings | 1 | Axial | 64 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 0 | 3.38 | 1000 | 13.51 | 1000 | 80
9.6 | nonHelical 8slice Routine - 80cm | Series 2 Group 2 Scan Settings | 1 | Axial | 64 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 0 | 3.38 | 1000 | 13.51 | 1000 | 80
9.6 | nonHelical 8slice Routine - 80cm | Series 2 Group 3 Scan Settings | 1 | Axial | 64 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 0 | 3.38 | 1000 | 13.51 | 1000 | 80
9.6 | nonHelical 8slice Routine - 80cm | Series 2 Group 4 Scan Settings | 1 | Axial | 64 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 0 | 3.38 | 1000 | 13.51 | 1000 | 80
9.6 | nonHelical 8slice Routine - 80cm | Series 2 Group 5 Scan Settings | 1 | Axial | 64 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 0 | 3.38 | 1000 | 13.51 | 1000 | 80
9.6 | nonHelical 8slice Routine - 80cm | Series 2 Group 6 Scan Settings | 1 | Axial | 64 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 0 | 3.38 | 1000 | 13.51 | 1000 | 80
9.6 | nonHelical 8slice Routine - 80cm | Series 2 Group 7 Scan Settings | 1 | Axial | 64 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 0 | 3.38 | 1000 | 13.51 | 1000 | 80
9.6 | nonHelical 8slice Routine - 80cm | Series 2 Group 8 Scan Settings | 1 | Axial | 64 | false | false | S0.0 | Large Body | 100 |  |  |  |  | 102 | 0 | 0 | 3.38 | 1000 | 13.51 | 1000 | 80
9.17 | nonHelical 8slice Routine - 80cm AEC | Series 2 Group 1 Scan Settings | 1 | Axial | 64 | false | false | S0.0 | Large Body | 100 | true | 10 | 100 | 700 | 102 | 0 | 0 | 8.45 | 1000 | 33.78 | 1000
9.17 | nonHelical 8slice Routine - 80cm AEC | Series 2 Group 2 Scan Settings | 1 | Axial | 64 | false | false | S0.0 | Large Body | 100 | true | 10 | 100 | 700 | 102 | 0 | 2 | 8.45 | 1000 | 33.78 | 1000
9.17 | nonHelical 8slice Routine - 80cm AEC | Series 2 Group 3 Scan Settings | 1 | Axial | 64 | false | false | S0.0 | Large Body | 100 | true | 10 | 100 | 700 | 102 | 0 | 2 | 8.45 | 1000 | 33.78 | 1000
9.17 | nonHelical 8slice Routine - 80cm AEC | Series 2 Group 4 Scan Settings | 1 | Axial | 64 | false | false | S0.0 | Large Body | 100 | true | 10 | 100 | 700 | 102 | 0 | 2 | 8.45 | 1000 | 33.78 | 1000
9.17 | nonHelical 8slice Routine - 80cm AEC | Series 2 Group 5 Scan Settings | 1 | Axial | 64 | false | false | S0.0 | Large Body | 100 | true | 10 | 100 | 700 | 102 | 0 | 2 | 8.45 | 1000 | 33.78 | 1000
9.17 | nonHelical 8slice Routine - 80cm AEC | Series 2 Group 6 Scan Settings | 1 | Axial | 64 | false | false | S0.0 | Large Body | 100 | true | 10 | 100 | 700 | 102 | 0 | 2 | 8.45 | 1000 | 33.78 | 1000
9.17 | nonHelical 8slice Routine - 80cm AEC | Series 2 Group 7 Scan Settings | 1 | Axial | 64 | false | false | S0.0 | Large Body | 100 | true | 10 | 100 | 700 | 102 | 0 | 2 | 8.45 | 1000 | 33.78 | 1000
9.17 | nonHelical 8slice Routine - 80cm AEC | Series 2 Group 8 Scan Settings | 1 | Axial | 64 | false | false | S0.0 | Large Body | 100 | true | 10 | 100 | 700 | 102 | 0 | 2 | 8.45 | 1000 | 33.78 | 1000
10.1 | GSI test 40mm Axial |  | 1 | Axial | 64 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 101 | 0 | 0 | 8.35 | 1000 | 33.4 | 1000 | 215
10.2 | GSI test 40mm Helical | Series 1 Group 1 Scan Settings | 0.5 | Helical | 64 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 101 | 0 | 0 | 9.31 | 1000 | 70.38 | 1000 | 200 | 33
10.3 | GSI test 40mm Helical HP | Series 1 Group 1 Scan Settings | 0.5 | Helical | 64 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 101 | 0 | 0 | 3.49 | 1000 | 27.19 | 1000 | 200 | 88
10.4 | GSI test 80mm Helical L | Series 1 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 101 | 0 | 0 | 9.24 | 1000 | 94.28 | 1000 | 200 | 65
10.5 | GSI test 80mm Helical L HP | Series 1 Group 1 Scan Settings | 1 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 101 | 0 | 0 | 6.35 | 1000 | 102.09 | 1000 | 185 | 176
10.6 | GSI test 80mm Helical XL | Series 1 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 101 | 0 | 0 | 26.13 | 1000 | 266.76 | 1000 | 530 | 65
10.7 | GSI test 80mm Helical XL HP | Series 1 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 101 | 0 | 0 | 9.65 | 1000 | 155.18 | 1000 | 530 | 176
10.8 | GSI test 40mm Head | Series 1 Group 1 Scan Settings | 1 | Helical | 64 | false | false | S0.0 | Adult Head | 120 |  |  |  |  | 101 | 0 | 0 | 57.31 | 1000 | 433.29 | 1000 | 270 | 33
10.9 | Smart Step | Series 2 Group 1 Scan Settings | 0.35 | Helical | 128 | false | false | S0.0 | Large Body | 120 | true | 14.7 | 150 | 700 | 103 | 0 | 2 | 12.73 | 1000 | 517.34 | 1000
10.9 | Smart Step | Series 3 Group 1 Scan Settings | 0.5 | Fluoro | 8 | false | false | S0.0 | Large Body | 80 | [7 empty] | 227 | 1000 | 113.5 | 1000 | 80
10.9 | Smart Step | Series 4 Group 1 Scan Settings | 0.35 | Helical | 128 | false | false | S0.0 | Large Body | 120 | true | 14.7 | 150 | 700 | 103 | 0 | 2 | 12.73 | 1000 | 517.34 | 1000
10.11 | Test Injection | Series 1 Group 1 Scan Settings | 1 | Axial | 8 | false | false | S0.0 | Large Body | 120 |  |  |  |  | 102 | 0 | 2 | 40.93 | 1000 | 20.47 | 1000 | 20
10.16 | Stress 50-59BPM | Series 2 Group 1 Scan Settings | 0.35 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Small | 120 | true | 32 | 50 | 200 | 102 | 0 | 1 | 1.09 | 1000 | 13.03 | 1000
10.16 | Stress 50-59BPM | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Small | 80 |  |  |  |  | 102 | 0 | 1 | 4.85 | 1000 | 48.55 | 1000 | 265
10.16 | Stress 50-59BPM | Series 4 Group 1 Scan Settings | 0.28 | Cardiac | 160 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 36.53 | 1000 | 365.25 | 1000 | 150
10.18 | Stress 70-89BPM | Series 2 Group 1 Scan Settings | 0.35 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Small | 120 | true | 32 | 50 | 200 | 102 | 0 | 1 | 1.09 | 1000 | 13.03 | 1000
10.18 | Stress 70-89BPM | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Small | 80 |  |  |  |  | 102 | 0 | 1 | 4.85 | 1000 | 48.55 | 1000 | 265
10.18 | Stress 70-89BPM | Series 4 Group 1 Scan Settings | 0.28 | Cardiac | 160 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 26.79 | 1000 | 267.85 | 1000 | 150
10.2 | Rest 30-39BPM | Series 1 Group 1 Scan Settings | 0.28 | Cardiac | 160 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 12.18 | 1000 | 121.75 | 1000 | 150
10.2 | Rest 30-39BPM | Series 1 Group 2 Scan Settings | 0.28 | Cardiac | 224 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 14.97 | 1000 | 209.58 | 1000 | 440
10.2 | Rest 30-39BPM | Series 1 Group 3 Scan Settings | 0.28 | Cardiac | 160 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 6.09 | 1000 | 60.88 | 1000 | 150
10.2 | Rest 30-39BPM | Series 2 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Small | 80 |  |  |  |  | 102 | 0 | 1 | 9.21 | 1000 | 92.1 | 1000 | 575
10.22 | Rest 50-59BPM | Series 1 Group 1 Scan Settings | 0.28 | Cardiac | 160 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 18.26 | 1000 | 182.63 | 1000 | 150
10.22 | Rest 50-59BPM | Series 1 Group 2 Scan Settings | 0.28 | Cardiac | 224 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 14.97 | 1000 | 209.58 | 1000 | 440
10.22 | Rest 50-59BPM | Series 1 Group 3 Scan Settings | 0.28 | Cardiac | 160 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 10.96 | 1000 | 109.58 | 1000 | 150
10.22 | Rest 50-59BPM | Series 2 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Small | 80 |  |  |  |  | 102 | 0 | 1 | 9.21 | 1000 | 92.1 | 1000 | 575
10.24 | Rest 70-89BPM | Series 1 Group 1 Scan Settings | 0.28 | Cardiac | 160 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 13.39 | 1000 | 133.93 | 1000 | 150
10.24 | Rest 70-89BPM | Series 1 Group 2 Scan Settings | 0.28 | Cardiac | 224 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 27.12 | 1000 | 379.73 | 1000 | 440
10.24 | Rest 70-89BPM | Series 1 Group 3 Scan Settings | 0.28 | Cardiac | 160 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 8.52 | 1000 | 85.23 | 1000 | 150
10.24 | Rest 70-89BPM | Series 2 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Small | 80 |  |  |  |  | 102 | 0 | 1 | 9.21 | 1000 | 92.1 | 1000 | 575
10.14 | Stress 30-39BPM | Series 2 Group 1 Scan Settings | 0.35 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Small | 120 | true | 32 | 50 | 200 | 102 | 0 | 1 | 1.09 | 1000 | 13.03 | 1000
10.14 | Stress 30-39BPM | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Small | 80 |  |  |  |  | 102 | 0 | 1 | 4.85 | 1000 | 48.55 | 1000 | 265
10.14 | Stress 30-39BPM | Series 4 Group 1 Scan Settings | 0.28 | Cardiac | 160 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 24.35 | 1000 | 243.5 | 1000 | 150
10.15 | Stress 40-49BPM | Series 2 Group 1 Scan Settings | 0.35 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Small | 120 | true | 32 | 50 | 200 | 102 | 0 | 1 | 1.09 | 1000 | 13.03 | 1000
10.15 | Stress 40-49BPM | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Small | 80 |  |  |  |  | 102 | 0 | 1 | 4.85 | 1000 | 48.55 | 1000 | 265
10.15 | Stress 40-49BPM | Series 4 Group 1 Scan Settings | 0.28 | Cardiac | 160 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 30.44 | 1000 | 304.38 | 1000 | 150
10.17 | Stress 60-69BPM | Series 2 Group 1 Scan Settings | 0.35 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Small | 120 | true | 32 | 50 | 200 | 102 | 0 | 1 | 1.09 | 1000 | 13.03 | 1000
10.17 | Stress 60-69BPM | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Small | 80 |  |  |  |  | 102 | 0 | 1 | 4.85 | 1000 | 48.55 | 1000 | 265
10.17 | Stress 60-69BPM | Series 4 Group 1 Scan Settings | 0.28 | Cardiac | 160 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 20.7 | 1000 | 206.98 | 1000 | 150
10.19 | Stress 90-140BPM | Series 2 Group 1 Scan Settings | 0.35 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Small | 120 | true | 32 | 50 | 200 | 102 | 0 | 1 | 1.09 | 1000 | 13.03 | 1000
10.19 | Stress 90-140BPM | Series 3 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Small | 80 |  |  |  |  | 102 | 0 | 1 | 4.85 | 1000 | 48.55 | 1000 | 265
10.19 | Stress 90-140BPM | Series 4 Group 1 Scan Settings | 0.28 | Cardiac | 160 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 28 | 1000 | 280.03 | 1000 | 150
10.21 | Rest 40-49BPM | Series 1 Group 1 Scan Settings | 0.28 | Cardiac | 160 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 14.61 | 1000 | 146.1 | 1000 | 150
10.21 | Rest 40-49BPM | Series 1 Group 2 Scan Settings | 0.28 | Cardiac | 224 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 14.97 | 1000 | 209.58 | 1000 | 440
10.21 | Rest 40-49BPM | Series 1 Group 3 Scan Settings | 0.28 | Cardiac | 160 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 9.74 | 1000 | 97.4 | 1000 | 150
10.21 | Rest 40-49BPM | Series 2 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Small | 80 |  |  |  |  | 102 | 0 | 1 | 9.21 | 1000 | 92.1 | 1000 | 575
10.23 | Rest 60-69BPM | Series 1 Group 1 Scan Settings | 0.28 | Cardiac | 160 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 10.96 | 1000 | 109.58 | 1000 | 150
10.23 | Rest 60-69BPM | Series 1 Group 2 Scan Settings | 0.28 | Cardiac | 224 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 27.12 | 1000 | 379.73 | 1000 | 440
10.23 | Rest 60-69BPM | Series 1 Group 3 Scan Settings | 0.28 | Cardiac | 160 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 7.31 | 1000 | 73.05 | 1000 | 150
10.23 | Rest 60-69BPM | Series 2 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Small | 80 |  |  |  |  | 102 | 0 | 1 | 9.21 | 1000 | 92.1 | 1000 | 575
10.25 | Rest 90-140BPM | Series 1 Group 1 Scan Settings | 0.28 | Cardiac | 160 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 13.39 | 1000 | 133.93 | 1000 | 150
10.25 | Rest 90-140BPM | Series 1 Group 2 Scan Settings | 0.28 | Cardiac | 224 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 20.85 | 1000 | 291.84 | 1000 | 440
10.25 | Rest 90-140BPM | Series 1 Group 3 Scan Settings | 0.28 | Cardiac | 160 | false | false | S0.0 | Cardiac Small | 100 |  |  |  |  | 102 | 0 | 0 | 12.18 | 1000 | 121.75 | 1000 | 150
10.25 | Rest 90-140BPM | Series 2 Group 1 Scan Settings | 0.28 | Cardiac | SmartCollimation | false | false | S0.0 | Cardiac Small | 80 |  |  |  |  | 102 | 0 | 1 | 9.21 | 1000 | 92.1 | 1000 | 575
11.5 | Child Head 6-10y Routine | Series 2 Group 1 Scan Settings | 0.7 | Axial | SmartCollimation | false | false | S0.0 | Small Head | 100 | false | 3.2 | 100 | 680 | 101 | 0 | 0 | 40.66 | 1000 | 569.3 | 1000
11.5 | Child Head 6-10y Routine | Series 3 Group 1 Scan Settings | 0.7 | Axial | SmartCollimation | false | false | S0.0 | Small Head | 100 | false | 3.2 | 100 | 680 | 101 | 0 | 0 | 34.48 | 1000 | 551.74 | 1000
11.3 | Child Head 0y Routine | Series 2 Group 1 Scan Settings | 0.6 | Axial | SmartCollimation | false | false | S0.0 | Small Head | 100 | false | 3.2 | 100 | 680 | 101 | 0 | 0 | 32.92 | 1000 | 460.86 | 1000
11.3 | Child Head 0y Routine | Series 3 Group 1 Scan Settings | 0.6 | Axial | SmartCollimation | false | false | S0.0 | Small Head | 100 | false | 3.2 | 100 | 680 | 101 | 0 | 0 | 34.6 | 1000 | 553.66 | 1000
11.4 | Child Head 1-5y Routine | Series 2 Group 1 Scan Settings | 0.8 | Axial | SmartCollimation | false | false | S0.0 | Small Head | 100 | false | 3.2 | 100 | 720 | 101 | 0 | 0 | 46.47 | 1000 | 650.63 | 1000
11.4 | Child Head 1-5y Routine | Series 3 Group 1 Scan Settings | 0.8 | Axial | SmartCollimation | false | false | S0.0 | Small Head | 100 | false | 3.2 | 100 | 720 | 101 | 0 | 0 | 34.6 | 1000 | 553.66 | 1000
11.2 | Child Brain non-Helical Routine ODM+ | Series 2 Group 1 Scan Settings | 1 | Axial | 256 | false | false | S0.0 | Adult Head | 100 | false | 4 | 200 | 400 | 101 | 0 | 0 | 36.88 | 1000 | 590.06 | 1000
11.1 | Child Brain-Head Routine ODM+ | Series 2 Group 1 Scan Settings | 0.8 | Helical | 64 | false | false | S0.0 | Small Head | 100 | false | 4 | 150 | 400 | 102 | 0 | 0 | 47.99 | 1000 | 1395.14 | 1000
12.1 | Child Orbit Sinus Helical Routine ODM+ | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Small Head | 100 | false | 10 | 100 | 330 | 101 | 0 | 0 | 8.12 | 1000 | 137.38 | 1000
12.1 | Child Orbit Sinus Helical Routine ODM+ | Series 3 Group 1 Scan Settings | 0.6 | Helical | 64 | false | false | S0.0 | Adult Head | 100 | true | 6.1 | 100 | 330 | 101 | 0 | 0 | 9.2 | 1000 | 166.24 | 1000
13.1 | Neck Volume | Series 2 Group 1 Scan Settings | 0.5 | Axial | SmartCollimation | false | false | S0.0 | Small Body | 100 | true | 10 | 100 | 500 | 101 | 0 | 0 | 7.39 | 1000 | 103.39 | 1000
13.2 | Child C-Spine Helical MAR Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Small Body | 100 | true | 12 | 150 | 700 | 102 | 0 | 2 | 8.69 | 1000 | 279.26 | 1000
14.1 | Child Shoulder Helical MAR | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | true | false | S0.0 | Small Body | 100 | true | 7.8 | 150 | 570 | 102 | 0 | 2 | 8.26 | 1000 | 237.14 | 1000
14.2 | Child Arm on Abd Helical MAR | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Small Body | 100 | true | 7.8 | 200 | 700 | 102 | 0 | 2 | 8.46 | 1000 | 237.87 | 1000
14.4 | Hand Tendon 140keV Helical | Series 2 Group 1 Scan Settings | 1 | Helical | 64 | true | false | S0.0 | Small Body | 140 | true | 3 | 200 | 405 | 102 | 0 | 0 | 16.46 | 1000 | 560.53 | 1000
14.3 | Child Arm over Head Helical MAR | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | true | false | S0.0 | Small Body | 100 | true | 7.8 | 100 | 450 | 102 | 0 | 0 | 8.66 | 1000 | 225.66 | 1000
15.1.5 | Child Arm 0y 1Volume Routine | Series 2 Group 1 Scan Settings | 0.35 | Axial | SmartCollimation | false | false | S0.0 | Small Body | 80 | false | 16.8 | 50 | 600 | 102 | 0 | 0 | 4.08 | 1000 | 65.31 | 1000
15.1.5 | Child Arm 0y 1Volume Routine | Series 3 Group 1 Scan Settings | 0.35 | Axial | SmartCollimation | false | false | S0.0 | Small Body | 80 | true | 16.8 | 50 | 600 | 102 | 0 | 0 | 4.08 | 1000 | 65.31 | 1000
15.1.6 | Child Arm1-5y Helical Routine | Series 2 Group 1 Scan Settings | 0.35 | Helical | 128 | false | false | S0.0 | Small Body | 100 | false | 16.8 | 50 | 700 | 102 | 0 | 0 | 6.77 | 1000 | 156.61 | 1000
15.1.6 | Child Arm1-5y Helical Routine | Series 3 Group 1 Scan Settings | 0.35 | Helical | 128 | false | false | S0.0 | Small Body | 100 | true | 16.8 | 50 | 700 | 102 | 0 | 0 | 6.77 | 1000 | 153.23 | 1000
15.1.4 | Child Chest 0y 1Volume Routine | Series 2 Group 1 Scan Settings | 0.35 | Axial | SmartCollimation | false | false | S0.0 | Small Body | 80 | false | 16.8 | 50 | 600 | 102 | 0 | 0 | 4.08 | 1000 | 65.31 | 1000
15.1.4 | Child Chest 0y 1Volume Routine | Series 3 Group 1 Scan Settings | 0.35 | Axial | SmartCollimation | false | false | S0.0 | Small Body | 80 | true | 16.8 | 50 | 600 | 102 | 0 | 0 | 4.08 | 1000 | 65.31 | 1000
15.1.1 | Child Chest 1-5y Helical Routine | Series 2 Group 1 Scan Settings | 0.35 | Helical | 128 | false | false | S0.0 | Small Body | 100 | false | 16.8 | 50 | 700 | 102 | 0 | 0 | 6.77 | 1000 | 156.61 | 1000
15.1.1 | Child Chest 1-5y Helical Routine | Series 3 Group 1 Scan Settings | 0.35 | Helical | 128 | false | false | S0.0 | Small Body | 100 | true | 16.8 | 50 | 700 | 102 | 0 | 0 | 6.77 | 1000 | 153.23 | 1000
15.1.2 | Child Chest 6-10y Helical Routine | Series 2 Group 1 Scan Settings | 0.28 | Helical | 128 | false | false | S0.0 | Small Body | 100 | false | 16.8 | 50 | 720 | 102 | 0 | 2 | 5.11 | 1000 | 125.59 | 1000
15.1.2 | Child Chest 6-10y Helical Routine | Series 3 Group 1 Scan Settings | 0.28 | Helical | 128 | false | false | S0.0 | Small Body | 100 | true | 16.8 | 50 | 720 | 102 | 0 | 2 | 5.11 | 1000 | 123.04 | 1000
15.1.3 | Child Arm 6-10y Helical Routine | Series 2 Group 1 Scan Settings | 0.28 | Helical | 128 | false | false | S0.0 | Small Body | 100 | false | 16.8 | 50 | 720 | 102 | 0 | 2 | 5.11 | 1000 | 125.59 | 1000
15.1.3 | Child Arm 6-10y Helical Routine | Series 3 Group 1 Scan Settings | 0.28 | Helical | 128 | false | false | S0.0 | Small Body | 100 | true | 16.8 | 50 | 720 | 102 | 0 | 2 | 5.11 | 1000 | 123.04 | 1000
15.1 | Child Chest Routine | Series 2 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Small Body | 100 | false | 14.7 | 100 | 700 | 102 | 0 | 2 | 8.83 | 1000 | 517.85 | 1000
15.1 | Child Chest Routine | Series 3 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Small Body | 100 | false | 14.7 | 100 | 700 | 102 | 0 | 2 | 8.83 | 1000 | 341.22 | 1000
16.1.1 | Child Abd 0y 1Volume Routine | Series 2 Group 1 Scan Settings | 0.5 | Axial | SmartCollimation | false | false | S0.0 | Small Body | 80 | false | 16.8 | 50 | 560 | 102 | 0 | 0 | 4.07 | 1000 | 65.19 | 1000
16.1.1 | Child Abd 0y 1Volume Routine | Series 3 Group 1 Scan Settings | 0.5 | Axial | SmartCollimation | false | false | S0.0 | Small Body | 80 | true | 16.8 | 50 | 560 | 102 | 0 | 0 | 4.07 | 1000 | 65.19 | 1000
16.1.2 | Child Abd 1-5y Helical Routine | Series 2 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Small Body | 100 | false | 16.8 | 50 | 560 | 102 | 0 | 0 | 6.78 | 1000 | 156.77 | 1000
16.1.2 | Child Abd 1-5y Helical Routine | Series 3 Group 1 Scan Settings | 0.5 | Helical | 128 | false | false | S0.0 | Small Body | 100 | true | 16.8 | 50 | 560 | 102 | 0 | 0 | 6.78 | 1000 | 153.39 | 1000
16.1.3 | Child Abd 6-10y Helical Routine | Series 2 Group 1 Scan Settings | 0.35 | Helical | 128 | false | false | S0.0 | Small Body | 120 | false | 16.8 | 50 | 600 | 102 | 0 | 2 | 10.43 | 1000 | 256.22 | 1000
16.1.3 | Child Abd 6-10y Helical Routine | Series 3 Group 1 Scan Settings | 0.35 | Helical | 128 | false | false | S0.0 | Small Body | 120 | true | 16.8 | 50 | 600 | 102 | 0 | 2 | 10.43 | 1000 | 251 | 1000
16.1 | Child Abdomen-Pelvis Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Small Body | 100 | false | 11.6 | 100 | 720 | 102 | 0 | 2 | 16.92 | 1000 | 864.87 | 1000
16.1 | Child Abdomen-Pelvis Routine | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Small Body | 100 | false | 11.6 | 100 | 720 | 102 | 0 | 2 | 16.92 | 1000 | 949.46 | 1000
17.1 | Child L-Spine Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Small Body | 100 | true | 11.6 | 150 | 700 | 102 | 0 | 2 | 8.46 | 1000 | 305.55 | 1000
18.1 | Child Pelvis Rotine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Small Body | 100 | false | 11.6 | 150 | 720 | 102 | 0 | 2 | 16.99 | 1000 | 793.82 | 1000
18.1 | Child Pelvis Rotine | Series 3 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Medium | 100 | true | 11.6 | 150 | 720 | 102 | 0 | 2 | 20.04 | 1000 | 924.41 | 1000
19.1 | Child Lower Extremity Heiical Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 128 | false | false | S0.0 | Small Body | 100 | true | 4.1 | 100 | 570 | 101 | 0 | 0 | 8.58 | 1000 | 481.33 | 1000
19.2 | Child Knee/Ankle Helical HR Routine | Series 2 Group 1 Scan Settings | 0.6 | Helical | 64 | true | false | S0.0 | Medium | 100 | true | 4.1 | 100 | 450 | 101 | 0 | 0 | 10.25 | 1000 | 267.11 | 1000